## Slide 1: Instructions for Presentation on Lab Procedures
Please use this slide deck from a TES training in Angola as a template for presenting on study laboratory procedures to your in-country TES team.
There are additional instructions in the notes for each slide.

## Slide 2: Procedures Laboratory

## Slide 3: DAILY ACTIVITIES
 | 0 | 1 | 2 | 3 | 7 | 14 | 21 | 28 |  |  | Ex
 | 0 | 1 | 2 | 3 | 7 | 14 | 21 | 28 |  |  | Ex
 | 0 | 1 | 2 | 3 | 7 | 14 | 21 | 28 | 35 | 42 | Ex
Parasitemia |  | Only serious malaria |  |  |  |  |  |  |  |  | 
Filter paper |  |  |  |  |  |  |  |  |  |  | 
Haemoglobin |  |  |  |  |  |  |  |  |  |  | 
AL
ASAQ
DP
Note: daily evaluation of the results will be done by the clinical staff involved in the study.
Ex: excluded
Clinical signs of severe malaria:
Changes in consciousness
Prostration
Multiple convulsions
Lab signs of severe malaria:
hyperparasitaemia (>4%)
renal insufficiency (elevated serum creatinine)
severe anaemia (haemoglobin < 5 g/dl)

## Slide 4: Three Responsibilities
1. microscopy
2. Haemoglobin measurement
3. Dried blood sample

## Slide 5: Finger prick
Putting on gloves
Disinfect of the patient's finger with alcohol
Leave to dry

## Slide 6: Finger Prick
4. Hold the finger firmly
5. Press the lancet on the finger

## Slide 7: Finger Prick
6. Press lightly on the finger to get 1 drop
7. Wipe the 1st drop from the finger

## Slide 8: Finger Prick
8. Prepare the microscope slide

## Slide 9: Finger Prick
9. Prepare the microcuvette

## Slide 10: Finger Prick
10. Prepare the filter paper

## Slide 11: Finger Prick
Put lancets in the sharps box and other material in the infectious waste

## Slide 12: 1: Microscopy

## Slide 13: Slide Preparation
(thick drop and smear)

## Slide 14: Labelling - Documentation
Each slide should receive a label, corresponding to the patient ID number
ZZ19-201-D1

## Slide 15: Slide Identification
Card No., (Number that corresponds to the patient’s study ID card)
Province and local code
The respective day (D0)
The date on which the sample was taken

## Slide 16: Slide Preparation

## Slide 17: Preparation of Giemsa stain 10% or 3%
Giemsa stain should be prepared (diluted) every day
  rest 30-60 mins before use
Dilute 10 cc or 3 cc of concentrated giemsa in a beaker
Add 90 or 97 cc of buffered water with pH = 7,2

## Slide 18: Preparing the slide
Dip the slide quickly in methyl alcohol and leave to dry
Place Giemsa stain on the slides
Leave the preparations in contact with the dye for 15-30 min

## Slide 19: Stained Slides

## Slide 20: Reading
Microscope with immersion lens using immersion oil.
Look for:
  Ringed forms (trophozoites)
  gametocytes with red nucleus or chromatin
  blue cytoplasm
Notes: *Fay will add

## Slide 21: Parasite count
The parasite count should be done from one end of the drop to the other.

## Slide 22: Parasite count
Parasites appearing in 200 white blood cells should be counted simultaneously.
For counting, it is recommended to use 2 cell counters
  One for the parasites
  Other for white blood cells (leukocytes)

## Slide 23: FORMULA
If you have counted 500 parasites before 200 white blood cells are counted you can stop counting after the last field is completed and do the math, using the following formula:
No. of parasites counted x 8000 leuc
No. of leuc counted
Parasites
by
microlitre
=

## Slide 24: 2 Readings
The slides should be recounted again at the end of the day
Confirmation of previous count
Internal quality control

## Slide 25: Discordance - 3rd Reading
____________________ – ____________________ = _________________
(higher parasitemia) (lower parasitemia) Difference
____________________ / ____________________ * 100 = ___________
Difference (lower parasitemia) %
Take 3rd reading if % is greater or equal to 25%.

## Slide 26: Discordance - 3rd Reading
40,000 – 30,000 = 10,000
(higher parasitemia) (lower parasitemia) Difference
10,000 / 30,000 * 100 = 33.3
Difference (lower parasitemia) %
1st Reading: 30,000
2nd Reading: 40,000
33.3% is greater than 25%
The 3rd reading is done

## Slide 27: Discordance - 3rd Reading
1st Reading: 65,000
2nd Reading: 75,000

## Slide 28: Discordance - 3rd Reading
75,000 – 65,000 = 10,000
(higher parasitemia) (lower parasitemia) Difference
10,000 / 65,000 * 100 = 15.4
Difference (lower parasitemia) %
1st Reading: 65,000
2nd Reading: 75,000
15.4% is NOT greater than 25%
NO 3rd reading

## Slide 29: Conservation of slides
The slide when drying should be:
  out of reach of insects and dust
  duly identified with number, code, date and day of collection
The number on the slide should be the same as the number on the study sheet
At the end of the reading, the slides should be carefully stored in appropriate boxes
  ATTN!!! Do not wash the blades with xylol

## Slide 30: Continued
After the slides have been read, they should be carefully placed, face down, on clean and suitable absorbent paper, to absorb the oil.
All the slides of the patients included in the study must be kept.
In the absence of appropriate boxes, these may be wrapped in absorbent paper, duly labelled and identified.

## Slide 31: Plasmodium sp

## Slide 32: Blood Smears

## Slide 33: Blood Smears

## Slide 34: Leukocytes

## Slide 35: Plasmodium vivax

## Slide 36: Plasmodium falciparum

## Slide 37: Plasmodium ovale

## Slide 38: PLASMODIUM MALARIAE

## Slide 39: 2: Haemoglobin

## Slide 40: HAEMOGLOBIN TEST OR HAEMATOCRIT
Haemoglobin is measured using a haemoglobinometer or a micro-hematocrit.

## Slide 41: Haemoglobin
Haemoglobin is measured in all study participants
Every 14 days
  D0, D14, D28 and D42
Also on unscheduled visits

## Slide 42: Haemoglobin I
Take a blood sample
Remove the cuvette from the cuvette container and cap the container immediately. Place the tip of the cuvette into the center of the drop of blood and allow the blood to collect in the cuvette
Fill the vial completely in a continuous process. It should never be completely full after the first filling (i.e. filled halfway, removed from near the drop of blood and then filled again)
If the cuvette is not completely filled with a drop of blood, take a new cuvette and fill it properly in a continuous process. NEVER use a half full cuvette for the sample

## Slide 43: Haemoglobin II
If air bubbles are present in the cuvette after filling with blood, use a new cuvette and obtain a new blood sample
Remove excess blood on the outside of the cuvette tip using a clean paper towel
In this process, make sure that there is no blood outside the cuvette

## Slide 44: Hemoglobin III - the Hemocue machine
Open the cuvette holder
Press the left button (the button with the circle) and the machine will switch on
An hourglass will appear in the lower left corner of the machine. This means that the machine is self-calibrating

## Slide 45: Hemoglobin IV- the Hemocue machine
When you see three flashing dashes and the HemoCue symbol, the machine is ready to take a measurement
Place the filled vial in the sample chamber of the HemoCue Photometer. The filled cuvette MUST be placed in the photometer within 10 minutes of sampling. Measurements taken after 10 minutes may result in erroneous haemoglobin counts
Insert the holder in the measuring position
During measurement, an hourglass will appear on the screen

## Slide 46: Haemoglobin V - the Hemocue machine
The results will appear within 15 to 60 seconds
Record the haemoglobin results
Remove the cuvette holder from the photometer, remove the cuvette that was read and dispose of it in the biological waste container
To turn the machine off, press and hold the left button until the display screen turns off

## Slide 47: Haemoglobin - Notes
NOTE: After 15 minutes of non-use the machine will switch off.
NOTE: Never press the power button for purposes other than turning the machine on or off. Pressing this button will show previous measurements and it will be difficult to tell which measurement is current and correct.

## Slide 48: Haemoglobin - Notes
Daily quality control
  The machine has its own quality control process. No control cuvettes are required. HemoCue machines should be cleaned every Friday with HemoCue cleaner in the health facility or whenever the study supervisor suspects that blood has been spilled inside the HemoCue machine

## Slide 49
Part 3: Filter Paper

## Slide 50: Filter Paper I
Materials required
  Protein saver blood collection card (Whatman 903 filter paper)
  Ziploc bags
  Desiccant pack
  Scissors
  Alcohol swab

## Slide 51: Filter Paper II
Open the protein saver blood collection card (Whatman 903 filter paper)
Label both ends of the card
Write the date on both labels
Collect blood onto filter paper ONLY after
  performing the haemoglobin test and
  preparing the blood smear

## Slide 52: Filter Paper III
The card contains 5 half-inch (1.27cm) circles that hold 50-100 μl each of the sample and a space for the name and date of the collection.

## Slide 53: Filter Paper IV
Label
Label

## Slide 54: Filter Paper V
Allow a drop of blood to form in the center of each spot shown or gently touch the drop of blood with the Whatman 903 blood collection card. Allow the drop of blood to evenly saturate the entire circle. (NB: do not add more than one drop per circle, do not touch the 903 Protein Holding Card at any time during collection; do not press the puncture site with the filter paper)
Fill in as many circles as possible on the card. Start at the edges and continue towards the middle.
After blood collection, ask the patient to press a small dry cotton ball on the puncture site for a few minutes to stop the bleeding.

## Slide 55: Filter Paper VI
Start at the extremities of the card so that both sides have at least one drop of blood

## Slide 56: Filter Paper VII
Start at the extremities of the card so that both sides have at least one drop of blood

## Slide 57: Filter Paper VIII
Start at the extremities of the card so that both sides have at least one drop of blood

## Slide 58: Filter Paper IX
Start at the extremities of the card so that both sides have at least one drop of blood

## Slide 59: Filter paper X
Start at the extremities of the card so that both sides have at least one drop of blood

## Slide 60: Filter paper XI
Let the 903 card dry:
  Away from direct sunlight
  in a clean place
  well ventilated and dry area
  protected against rodents and insects
Place the 903 card on a bench
  for at least 4 hours until the drops of blood have completely dried
  but not > 24 hours at room temperature

## Slide 61: Filter Paper XII
When dry
cut the card in half with clean scissors (sterilize using an alcohol swab)
make sure that both cut halves of the 903 card contain the child's identification number and date.

## Slide 62: Filter Paper XIII
Close the cover of the two halves of the 903 card
place each one in an individual Ziploc bag
place a desiccant packet inside each bag and seal

## Slide 63: Filter Paper XIII
The sealed and dry 903 protein saver card can be stored at room temperature (if the room temperature does not exceed 30 °C) for up to 3 weeks.
Afterwards, the closed and dried 903 cards can be stored at 4 °C for a period not exceeding 4 months.
For long-term storage, cards should be stored in the freezer at -70 °C for the first 3 weeks after sample collection.

## Slide 64: Examples of Poor Quality Samples
A. Insufficient quantity of blood
B. Too much blood
C. Samples that appear coagulated or clumped
D.Samples that appear diluted, discoloured or contaminated
E. Samples with alcohol circles
F. Samples not dried before storage